DOCUMENT 4

# Clinical Evidence Specification

PART 11

# Clinical Validation, Limitations, and Local Applicability

Verification, Model Evaluation, Workflow Testing, Philippine Lying-in Context,
Human Factors, Equity, Monitoring, and Deployment Governance

Purpose of this chapter Part 11 defines what must be validated before the CDSS can be considered safe for supervised use, what the current evidence does and does not establish, and how the system must be adapted to the realities of a Philippine lying-in clinic. It separates technical correctness from clinical usefulness and makes clear that successful classroom or test execution is not equivalent to demonstrated patient benefit.

Capstone Project

Development of an Explainable Clinical Decision Support System (CDSS)
for a Web-Based Maternity Management System in a Philippine Lying-in Clinic

Rewritten evidence specification | July 2026

# Executive Summary

Clinical validation is not a single accuracy number. For this hybrid system it includes verification of deterministic rules, validation of the Random Forest model, evaluation of the combined human-CDSS workflow, assessment of recommendation and explanation safety, usability testing with intended users, confirmation of referral feasibility, subgroup and equity analysis, and real-world monitoring after deployment.

IMDRF describes clinical evaluation of software as an ongoing process that considers valid clinical association, analytical or technical validation, and clinical validation. NICE requires evidence of effectiveness, real-world performance, monitoring over time, deployment transparency, user communication, and training. DECIDE-AI emphasizes early live clinical evaluation of actual performance, safety, workflow, and human factors before larger trials. WHO requires protection of human autonomy, safety, transparency, accountability, inclusiveness, and responsiveness throughout the AI lifecycle.

Current evidence status The existing capstone implementation and automated tests can demonstrate software verification and internal logical consistency. They do not, by themselves, establish external clinical validity, improvement in maternal outcomes, safety in unsupervised use, regulatory approval, or generalizability to other clinics. The system should therefore be presented as a supervised research prototype pending prospective local evaluation.

# 1. Validation Concepts and Boundaries

| Term | Meaning for this project | Not equivalent to |
| --- | --- | --- |
| Software verification | The code implements the documented requirements and handles expected inputs and errors. | Clinical benefit or correct diagnosis. |
| Rule validation | Thresholds, precedence, explanations, and recommendations match the approved evidence specification. | Proof that every patient classified HIGH will have a complication. |
| Model validation | The Random Forest is evaluated on data not used to fit it, using appropriate discrimination, calibration, error, and subgroup measures. | Generalizability to a Philippine lying-in clinic without local data. |
| Clinical validation | Evidence that the output is meaningfully associated with the intended clinical state or outcome and performs acceptably when used as intended. | A single retrospective test split. |
| Clinical utility | Evidence that using the CDSS improves decisions, referrals, timeliness, workflow, or patient outcomes compared with usual care. | High AUC alone. |
| Usability validation | Intended users can understand and safely act on the system under realistic conditions. | A developer demonstration. |
| Local applicability | The population, data, workflow, language, facility capability, and referral network match the deployment setting. | Use of international guidelines alone. |
| Post-deployment monitoring | Ongoing measurement of use, failures, drift, subgroup performance, overrides, and safety events. | One-time acceptance testing. |

# 2. Validation Framework

| Domain | Primary question | Evidence needed |
| --- | --- | --- |
| Clinical association | Do the inputs and outputs have a defensible relationship to the intended risk or care need? | Parts 1-7, clinician review, outcome definition. |
| Analytical/technical performance | Does the software correctly process inputs and produce reproducible outputs? | Unit, integration, security, reliability, and boundary tests. |
| Model performance | How well does the Random Forest discriminate and calibrate on unseen data? | Locked test set, confidence intervals, calibration, subgroup analysis. |
| Human factors | Can midwives understand, verify, and act without harmful overreliance? | Simulation, usability sessions, task completion, comprehension, error analysis. |
| Workflow effectiveness | Does the system improve completeness, referral, timeliness, or documentation? | Prospective silent/live pilot compared with usual workflow. |
| Safety | What harms, delays, false reassurance, over-referral, or privacy risks occur? | Hazard analysis, incident monitoring, stopping rules. |
| Equity | Does performance or access differ by age, socioeconomic context, geography, or other relevant groups? | Pre-specified subgroup and barrier analysis. |
| Local deployability | Can the clinic and referral network carry out the recommendation? | Capability mapping, transport, contacts, operating hours, acceptance process. |
| Lifecycle performance | Does performance remain stable after data, software, workflow, or population changes? | Versioning, drift monitoring, periodic revalidation. |

# 3. Deterministic Rule Verification

Deterministic rules are directly testable and should achieve exact conformance with the approved specification. Every threshold boundary, OR condition, missing-data state, precedence rule, interaction, recommendation, and explanation must have automated tests.

| Verification area | Required tests | Acceptance criterion |
| --- | --- | --- |
| Age | 18/19 boundary; 34/35 boundary; first-pregnancy true/false; invalid age. | Exact documented behavior; no hidden rule. |
| Blood pressure | 139/89, 140/80, 120/90, 159/109, 160/80, 120/110; missing component; repeat value. | OR logic correct; urgent precedence correct. |
| Medical history | Diabetes, anemia, previous CS true/false/unknown; string/boolean normalization. | No blank-to-No conversion. |
| Miscarriage | Counts 0,1,2,3; negative, decimal, duplicate, nonnumeric. | Integer validation; threshold >=3. |
| Ultrasound | All controlled values, case normalization, unknown, missing, contradictory, stale report. | No unsupported value or automatic normal. |
| Completeness | Each required record absent alone and in combination; HIGH plus missing data. | LOW prohibited; verified HIGH retained. |
| ML interface | LOW, HIGH, invalid, error text, timeout, missing executable, malformed response. | Only exact valid output accepted. |
| Interactions | Every Part 8 interaction and multi-factor cluster. | No undocumented interaction or duplicated reason. |
| Recommendations | Every Part 9 package and referral status. | Correct urgency and no prohibited treatment. |
| Explanations | Every Part 10 pathway and data-uncertainty message. | Faithful to actual decision path. |
| Delivered workflow | Read-only enforcement and new-pregnancy separation. | Historical integrity preserved. |
| Audit | Create/update/recalculate/override/refusal events. | Complete immutable trace. |

Rule acceptance target For deterministic requirements, the expected target is not approximate accuracy. Every approved test must pass. Any mismatch between specification, code, interface, referral printout, or explanation is a release-blocking defect.

# 4. Random Forest Model Validation

## 4.1 Dataset and outcome audit

Define the prediction target precisely, including clinical meaning, reference standard, assessment time, and who assigned the label.

Document data source, collection dates, institutions, inclusion/exclusion criteria, sample size, missingness, duplicates, and class balance.

Prevent patient-level leakage across training and test sets, including repeated pregnancies or visits from the same patient.

Confirm that every feature was available at the intended prediction time and is represented consistently in the Laravel workflow.

Compare the development population with the intended Philippine lying-in population.

Investigate label quality and whether historical practice patterns or referral decisions introduced bias.

## 4.2 Internal validation design

| Method | Appropriate use | Key safeguard |
| --- | --- | --- |
| Train/validation/test split | Adequate dataset with a truly untouched final test set. | Split by patient, not row or visit. |
| Cross-validation | Model selection and uncertainty estimation when data are limited. | All preprocessing occurs inside each fold. |
| Bootstrap | Performance optimism and confidence intervals. | Preserve patient grouping. |
| Temporal validation | Later time period from same source. | Tests change over time. |
| Geographic/external validation | Different clinic, province, or health system. | Required before broad claims of generalizability. |

## 4.3 Required performance measures

| Measure | Why needed | Caution |
| --- | --- | --- |
| Sensitivity/recall for HIGH | Measures how many reference HIGH cases the model identifies. | Threshold and class balance affect result. |
| Specificity | Measures avoidance of false HIGH among reference LOW cases. | High specificity cannot compensate for unsafe false negatives. |
| Positive/negative predictive value | Reflects practical meaning in the tested prevalence. | Changes when prevalence changes. |
| ROC AUC | Measures ranking discrimination across thresholds. | Does not show calibration or clinical usefulness. |
| Precision-recall AUC | Useful when HIGH is uncommon. | Must report baseline prevalence. |
| Calibration plot/intercept/slope | Assesses agreement between predicted and observed risk. | Required before displaying probabilities. |
| Confusion matrix | Shows actual error counts. | Report with denominators. |
| Confidence intervals | Shows statistical uncertainty. | Small samples can make estimates unstable. |
| Decision-curve or utility analysis | Explores net benefit across thresholds when applicable. | Requires defensible consequences and threshold interpretation. |
| Subgroup performance | Detects potentially unequal errors. | Avoid declaring fairness from underpowered comparisons. |

Categorical output limitation Because the current application uses only LOW/HIGH and may not display a probability, the thesis must still document how the classification threshold was selected. A high test accuracy without calibration, error analysis, and an external dataset is insufficient for a clinical claim.

# 5. Hybrid System Validation

| Scenario | Required evaluation |
| --- | --- |
| Rule HIGH, ML LOW | Confirm deterministic override, explanation fidelity, urgent/referral package, and audit. |
| Rule LOW conditions, ML HIGH | Confirm final HIGH is identified as ML-only and reviewed safely. |
| Rule HIGH, ML error | Confirm HIGH remains available while technical error is visible. |
| No rule, ML error | Confirm ASSESSMENT INCOMPLETE and manual pathway. |
| Multiple rules | Confirm all reasons retained, correct urgency, no unsupported score. |
| Missing record plus HIGH | Confirm HIGH retained and data gaps shown. |
| Complete LOW | Confirm all completeness conditions before LOW. |
| Model unavailable for prolonged period | Confirm deterministic/manual workflow remains usable and staff know limitations. |
| Model version change | Confirm locked validation, release approval, migration, and historical version preservation. |
| Rule/model disagreement rate | Measure frequency, clinical review, and whether model adds useful information. |

# 6. Clinical Workflow and Human-Factors Evaluation

DECIDE-AI emphasizes early-stage live evaluation of actual clinical performance, safety, and human-computer interaction. A capstone pilot should therefore evaluate the combined midwife-system team, not only the algorithm.

| Evaluation task | Suggested measure |
| --- | --- |
| Correct patient selection | Percentage of assessments linked to the correct pregnancy episode. |
| Data entry | Completion rate, error rate, correction rate, and time per record. |
| Risk recognition | Percentage correctly identifying HIGH, urgent, LOW, and incomplete states. |
| Explanation comprehension | Ability to state the reason, uncertainty, and next action without prompting. |
| Referral task | Correct facility capability, referral creation time, acceptance and completion. |
| Emergency simulation | Time to recognize and act on severe BP or acute danger scenario. |
| Automation bias | Whether users follow an incorrect or intentionally discordant test recommendation. |
| Alert fatigue | Dismissals, ignored alerts, duplicate alerts, and perceived burden. |
| Patient communication | Comprehension, emotional response, stigma, and trust. |
| Recovery from failure | Ability to continue safely during ML error, network failure, or missing record. |
| Training adequacy | Performance before/after training and retention over time. |
| Usability score | Validated tool such as SUS plus task-based qualitative findings. |

# 7. Prospective Local Evaluation Plan

| Phase | System exposure | Primary purpose | Minimum release condition |
| --- | --- | --- | --- |
| Phase 0: Document and code verification | Synthetic/test data only. | Confirm logic, security, and traceability. | All critical tests pass. |
| Phase 1: Retrospective local validation | Historical de-identified records; no effect on care. | Assess model/rule behavior and data compatibility. | Pre-specified performance and error review. |
| Phase 2: Silent prospective pilot | System runs on current cases but output hidden from clinical decisions. | Measure missingness, drift, disagreement, timing, and feasibility. | No unacceptable safety/data-quality issue. |
| Phase 3: Supervised live pilot | Output visible to trained staff; clinician oversight; limited site and duration. | Assess human-CDSS performance, referrals, usability, and safety. | Governance review against stopping rules. |
| Phase 4: Comparative evaluation | Compare against usual care or phased implementation. | Estimate effect on completeness, referral, timeliness, workflow, and patient outcomes. | Study protocol and ethics approval. |
| Phase 5: Controlled scale-up | Additional sites after local adaptation and external validation. | Assess transportability and equity. | Site-specific readiness and monitoring. |

Research ethics requirement Any study using identifiable patient records or exposing live clinical decisions to the CDSS requires appropriate institutional, clinical, privacy, and research-ethics review. Patient consent or another lawful basis must be determined by the responsible institution; the capstone team should not assume that ordinary clinic registration automatically authorizes research use.

# 8. Safety Evaluation and Stopping Rules

| Hazard | Example indicator | Suggested action |
| --- | --- | --- |
| False LOW | Reference/high-risk case classified LOW or missing rule. | Immediate case review; suspend affected release if logic defect. |
| Delayed urgent transfer | Urgent alert not recognized or referral delayed. | Immediate incident review and workflow correction. |
| Over-referral | Large increase in stable referrals without clinical benefit. | Review thresholds, wording, and local pathway capacity. |
| Data mismatch | Wrong pregnancy, stale ultrasound, incorrect BP units, duplicate history. | Correct data and investigate system safeguards. |
| Explanation mismatch | Displayed reason does not match active rule. | Release-blocking defect. |
| Privacy breach | Unauthorized access, disclosure, or export. | Activate privacy incident response. |
| Model drift | Performance or prevalence changes beyond pre-specified range. | Investigate and suspend/revalidate if necessary. |
| Subgroup harm | Materially worse false-negative or referral performance. | Clinical/equity review and corrective action. |
| Referral-network failure | Receiving facility cannot accept recommended case. | Update capability registry and escalation alternatives. |
| User overreliance | Staff treats CDSS as diagnosis or ignores symptoms. | Retraining, interface change, and governance review. |

# 9. Local Applicability to a Philippine Lying-in Clinic

Philippine birthing homes and lying-in clinics are non-hospital facilities centered on antenatal and postnatal care, normal spontaneous delivery, and newborn services. PhilHealth describes normal spontaneous delivery packages as intended for normal, low-risk vaginal deliveries, while cesarean and complicated deliveries are hospital services. This makes referral capability, transport, and continuity of care central to local validity.

| Local characteristic | Validation implication |
| --- | --- |
| Midwife-led workflow | Language, task order, and scope must be tested with actual midwives rather than assumed from hospital workflows. |
| Limited diagnostics | Recommendations must account for external laboratory and ultrasound services and report delays. |
| No operating theatre/blood bank | Previous cesarean, severe hypertension, malpresentation, and other complex cases require realistic hospital pathways. |
| Variable connectivity/equipment | The system needs safe failure behavior, local backups, and a plan for unavailable Python service or internet. |
| Referral delays and transport barriers | Measure acceptance, travel time, cost, transport availability, and failed referrals. |
| Multiple record sources | Validate manual transcription, paper-to-digital transfer, and duplicate-pregnancy prevention. |
| Language and health literacy | Test English and any Filipino/local-language counseling for meaning and urgency. |
| Resource constraints | Avoid recommendations that depend on unavailable specialists, tests, or medications without an alternative pathway. |
| Continuity across pregnancy | Ensure referral feedback, delivered status, and new-pregnancy linkage work in practice. |
| Privacy in small communities | Use role-based access, discreet communication, and minimum necessary disclosure. |

# 10. Population Transportability and Equity

| Issue | Risk | Required evaluation |
| --- | --- | --- |
| Development data from another country/setting | Different disease prevalence, care pathways, demographics, and labels. | External Philippine validation before clinical performance claims. |
| Adolescent patients | Consent, privacy, social vulnerability, and different baseline risk. | Age-specific usability and referral-barrier review. |
| Older first-pregnancy patients | Small subgroup may produce unstable estimates. | Report counts and uncertainty. |
| Low-income/rural patients | Referral may be less feasible despite identical recommendation. | Measure completion and barriers, not only recommendation generation. |
| Missing-data patterns | Patients with fewer resources may have more incomplete records. | Compare incomplete rates and false reassurance by subgroup. |
| Language/education | Explanation comprehension may differ. | User testing and clinician-reviewed translation. |
| Rare conditions | Performance estimates may be unreliable. | Do not overstate accuracy; use rule safeguards and specialist review. |
| Facility variation | A recommendation feasible at one clinic may fail elsewhere. | Site-specific capability assessment. |
| Historical bias | Labels may reflect previous access and clinician decisions. | Audit target definition and proxy variables. |
| Digital access/privacy | Some patients may be disadvantaged by email or electronic follow-up. | Provide non-digital alternatives. |

# 11. Key Clinical Limitations

| Limitation | Consequence |
| --- | --- |
| The CDSS is not a diagnostic device | HIGH and LOW are decision-support categories, not confirmed diagnoses. |
| Only HIGH, LOW, and INCOMPLETE are implemented | Clinical severity cannot always be represented by three categories. |
| Deterministic rules are conservative project configurations | They may over-refer and do not replace condition-specific diagnosis. |
| Several fields are boolean or broad categories | Severity, timing, cause, treatment response, and context are lost. |
| Ultrasound is transcribed, not interpreted | The system cannot validate image quality or diagnose findings. |
| BP may be based on one visit entry | Measurement error and transient elevation are possible. |
| ML outcome definition and dataset may be limited | Model performance may not transfer to the local population. |
| No validated probability is shown | The categories do not quantify absolute risk. |
| Interactions are recommendation logic, not a validated clinical score | Multiple factors do not produce a measured combined probability. |
| Referral recommendation does not guarantee access | Transport, cost, acceptance, and patient preference remain outside the algorithm. |
| A midwife remains the primary local user | System advice must remain within scope and require physician/hospital escalation. |
| No evidence yet of improved maternal or neonatal outcomes | Prospective comparative study is required. |

# 12. Technical and Operational Limitations

| Limitation | Control or future requirement |
| --- | --- |
| Python process/path dependency | Health check, clear error state, safe deterministic fallback, deployment documentation. |
| Local machine configuration | Reproducible environment, dependency lock, backups, and setup validation. |
| Database quality | Validation constraints, pregnancy linkage, deduplication, audit, and correction workflow. |
| Manual data entry | Controlled fields, units, range checks, source/date fields, and staff training. |
| Email-only notifications | Do not treat email as emergency communication; provide phone/transfer workflow. |
| Limited interoperability | Use structured exports/referral summaries and future standards mapping. |
| Rule duplication across code/views | Centralize logic and shared identifiers. |
| Model updates | Locked versions, validation, approval, rollback, and historical trace. |
| Cybersecurity | Authentication, least privilege, patching, encryption, backup, logging, and incident response. |
| System downtime | Paper/manual contingency and later reconciliation. |
| Single-site testing | External site validation before scale-up. |
| Small sample size | Report uncertainty and avoid broad claims. |

# 13. Privacy, Consent, and Data Governance

| Requirement | Local implementation |
| --- | --- |
| Lawful and transparent processing | Document purpose, legal basis, data categories, recipients, retention, and patient rights. |
| Automated-decision information | Explain rule and ML use, significance, envisaged consequences, and human review. |
| Data minimization | Collect only fields needed for care, validation, and approved research. |
| Role-based access | Restrict clinical, administrative, developer, and research access. |
| Research dataset controls | De-identify/pseudonymize where appropriate and separate from live care database. |
| Access/correction | Provide workflow for patients and authorized staff to correct inaccurate records. |
| Retention/deletion | Use approved clinic and research retention schedules. |
| Data sharing | Use agreements and minimum necessary disclosure with referral and research partners. |
| Security monitoring | Log access, exports, failed authentication, changes, and incidents. |
| Secondary use | Do not reuse data for model training or publication without institutional authorization. |

# 14. Monitoring After Deployment

| Metric group | Examples |
| --- | --- |
| Use | Assessments per month, eligible patients assessed, completion time, abandonment. |
| Data quality | Missing fields, invalid values, stale reports, duplicates, correction rate. |
| Rule behavior | HIGH/LOW/incomplete rates, active rules, multi-factor cases, disagreement with ML. |
| Model performance | Sensitivity, specificity, predictive values, calibration if probabilities exist, drift. |
| Subgroups | Performance and incomplete/referral rates by relevant age, location, and access groups. |
| Safety | False LOW reviews, delayed urgent action, incidents, near misses, overrides. |
| Referral | Created, accepted, completed, refused, unsuccessful; time to completion. |
| Human factors | Alert acknowledgement, comprehension, workarounds, alert fatigue. |
| Patient impact | Understanding, satisfaction, barriers, anxiety or stigma complaints. |
| Operations | Downtime, Python failures, latency, backup restore, security events. |
| Outcomes | Where ethically and feasibly collected: complications, transfer, delivery facility, maternal/newborn outcomes. |

# 15. Acceptance Criteria for a Supervised Pilot

| Area | Minimum criterion before pilot |
| --- | --- |
| Clinical documentation | Parts 1-12 reviewed and approved by qualified clinical governance. |
| Software | All critical deterministic, workflow, security, and audit tests pass. |
| Model | Locked version with transparent dataset and internal validation report; limitations accepted. |
| Data | Local field mapping and data-quality pilot completed. |
| Usability | Midwives successfully complete critical tasks and recognize emergency states. |
| Referral | Receiving facilities, contacts, transport, and refusal workflow confirmed. |
| Safety | Hazard analysis, incident pathway, stopping rules, and manual fallback approved. |
| Privacy/ethics | Privacy notice, access controls, institutional approval, and research ethics as applicable. |
| Training | Users trained on scope, limitations, urgent pathways, and downtime procedures. |
| Monitoring | Pre-specified metrics, reviewer, review frequency, and escalation thresholds. |
| Patient communication | Clinically reviewed patient wording and language process. |
| Governance | Named clinical owner, technical owner, privacy owner, and deployment decision authority. |

# 16. Validation Test Matrix

| Test ID | Validation scenario | Required result |
| --- | --- | --- |
| VAL-T01 | All documented deterministic boundary tests | 100% pass. |
| VAL-T02 | Controller/view/referral explanation consistency | Same rule IDs, result, urgency, and wording source. |
| VAL-T03 | Patient-level train/test leakage audit | No patient or pregnancy leakage. |
| VAL-T04 | Locked test set performance | Report all pre-specified metrics and CIs. |
| VAL-T05 | Calibration assessment | Reported; probabilities hidden if not sufficiently validated. |
| VAL-T06 | Subgroup analysis | Counts, uncertainty, and clinically reviewed differences. |
| VAL-T07 | Temporal or external validation | Completed before broad generalizability claim. |
| VAL-T08 | ML unavailable simulation | Safe deterministic/manual workflow. |
| VAL-T09 | Severe BP simulation | User recognizes and initiates urgent pathway. |
| VAL-T10 | Absent cardiac activity scenario | No auto-diagnosis; correct confirmation/referral wording. |
| VAL-T11 | Wrong patient/pregnancy prevention | User cannot silently attach record to wrong pregnancy. |
| VAL-T12 | Stale ultrasound | Date warning and verification task displayed. |
| VAL-T13 | Duplicate miscarriage record | Detected or corrected before threshold use. |
| VAL-T14 | HIGH plus missing records | HIGH retained; gaps displayed. |
| VAL-T15 | No HIGH plus missing records | INCOMPLETE, never LOW. |
| VAL-T16 | Alert-fatigue session | No critical alert missed; nonurgent burden acceptable. |
| VAL-T17 | Automation-bias test | User challenges intentionally discordant output. |
| VAL-T18 | Patient explanation comprehension | Patient can state meaning and next action. |
| VAL-T19 | Referral acceptance workflow | Status reflects real handover, not only print action. |
| VAL-T20 | Referral refusal | HIGH retained and nonjudgmental safety-net documentation. |
| VAL-T21 | Offline/downtime workflow | Care continues and records reconcile safely. |
| VAL-T22 | Access-control test | Unauthorized users cannot view or change sensitive data. |
| VAL-T23 | Audit reconstruction | Independent reviewer reconstructs decision from log. |
| VAL-T24 | New model/version release | Validation, approval, rollback, and historical version preserved. |
| VAL-T25 | Delivered/new pregnancy workflow | Historical/current data separated. |
| VAL-T26 | Local-language review | Meaning and urgency equivalent to approved English. |
| VAL-T27 | Receiving-facility capability check | Recommended services actually available. |
| VAL-T28 | Prospective silent pilot | Pre-specified data quality and safety thresholds met. |
| VAL-T29 | Supervised live pilot | No stopping-rule breach; workflow benefit or feasibility demonstrated. |
| VAL-T30 | Post-deployment monitoring drill | Drift/safety signal detected, reviewed, and acted upon. |

# 17. Panel Defense Questions and Model Answers

Q: Is your system clinically validated?

A: It has software and logic verification, but it should not be claimed as fully clinically validated until the Random Forest is externally evaluated and the complete human-CDSS workflow is prospectively tested in the intended Philippine setting.

Q: Why is accuracy not enough?

A: Accuracy can hide class imbalance, poor sensitivity, miscalibration, subgroup harm, workflow failures, and unsafe false reassurance.

Q: What must be validated separately?

A: The deterministic rules, the ML model, their hybrid precedence, explanations, recommendations, usability, referral workflow, security, and local deployability.

Q: Why do you need local validation?

A: Population, prevalence, data quality, language, staff workflow, resources, and referral access may differ from the development setting.

Q: What is the role of DECIDE-AI?

A: It guides reporting of early live clinical evaluation, including actual performance, safety, workflow, and human factors.

Q: What would be your safest first clinical study?

A: A silent prospective pilot in which the system runs on current cases but does not influence care, followed by a supervised live pilot after safety review.

Q: Can you say the model improves maternal outcomes?

A: No. That requires a prospective comparative evaluation with appropriate outcome measurement.

Q: How do you handle bias?

A: Audit dataset and labels, report subgroup performance and missingness, test explanation comprehension and referral completion, and monitor inequitable errors or access barriers.

Q: What is the main local limitation?

A: A recommendation is only useful if a real hospital, transport pathway, and qualified clinician are accessible. The software cannot create missing health-system capacity.

Q: When should deployment stop?

A: Examples include a false-LOW logic defect, delayed emergency action, explanation mismatch, privacy breach, serious subgroup harm, or model drift beyond pre-specified limits.

# 18. Final Validation and Applicability Statement

Approved specification The CDSS shall be described as a supervised research prototype until its deterministic logic, Random Forest model, hybrid decision pathway, explanations, recommendations, human factors, referral workflow, privacy controls, and local deployment assumptions have been evaluated through a staged validation programme. Deterministic rules require exact conformance testing. The model requires transparent patient-level internal validation, calibration and error analysis, subgroup evaluation, and external or temporal validation before broad clinical claims. Prospective local evaluation shall begin with silent observation and progress to limited supervised use only after clinical, technical, privacy, ethics, and referral-readiness approval. Validation shall measure the performance of the human-CDSS team and real referral process, not only algorithm accuracy. The system shall not be claimed to diagnose disease, improve maternal outcomes, generalize to other settings, or operate safely without human oversight unless evidence specifically supports those claims. Performance, safety, equity, data quality, user behavior, referrals, and technical reliability shall be monitored throughout the lifecycle.

# 19. References

Official regulatory, guideline, and primary reporting sources are used wherever publicly available. Accessed July 2026. URLs are clickable.

[1] International Medical Device Regulators Forum. Software as a Medical Device: Clinical Evaluation. 2017. https://www.imdrf.org/documents/software-medical-device-samd-clinical-evaluation

[2] International Medical Device Regulators Forum. Good Machine Learning Practice for Medical Device Development: Guiding Principles. 2025. https://www.imdrf.org/documents/good-machine-learning-practice-medical-device-development-guiding-principles

[3] U.S. Food and Drug Administration. Good Machine Learning Practice for Medical Device Development: Guiding Principles. 2025. https://www.fda.gov/medical-devices/software-medical-device-samd/good-machine-learning-practice-medical-device-development-guiding-principles

[4] World Health Organization. Ethics and governance of artificial intelligence for health. 2021. https://www.who.int/publications/i/item/9789240029200

[5] World Health Organization. Recommendations on digital interventions for health system strengthening. 2019. https://www.who.int/publications/i/item/9789241550505

[6] National Institute for Health and Care Excellence. Evidence standards framework for digital health technologies (ECD7). 2018; updated 2022. https://www.nice.org.uk/corporate/ecd7

[7] National Institute for Health and Care Excellence. Evidence standards framework: evidence standards tables. 2022. https://www.nice.org.uk/corporate/ecd7/chapter/section-c-evidence-standards-tables

[8] Vasey B, et al.. DECIDE-AI: reporting guideline for early-stage clinical evaluation of AI decision support systems. 2022. https://www.nature.com/articles/s41591-022-01772-9

[9] Collins GS, et al.. TRIPOD+AI statement: updated guidance for reporting clinical prediction models. 2024. https://www.bmj.com/content/385/bmj-2023-078378

[10] U.S. Food and Drug Administration, Health Canada, MHRA. Transparency for Machine Learning-Enabled Medical Devices: Guiding Principles. current online edition. https://www.fda.gov/medical-devices/software-medical-device-samd/transparency-machine-learning-enabled-medical-devices-guiding-principles

[11] National Privacy Commission, Philippines. The right to be informed. current online guidance. https://privacy.gov.ph/the-right-to-be-informed/

[12] National Privacy Commission, Philippines. Implementing Rules and Regulations of the Data Privacy Act of 2012. 2016. https://privacy.gov.ph/implementing-rules-regulations-data-privacy-act-2012/

[13] Department of Health, Philippines. Administrative Order No. 2007-0039: Regulation of Birthing Homes. 2007. https://elibrary.judiciary.gov.ph/thebookshelf/showdocs/10/52132

[14] Philippine Health Insurance Corporation. Women about to give birth may avail themselves of PhilHealth benefits. 2016. https://www.philhealth.gov.ph/news/2016/pregnancy_benefits.html

[15] Philippine Health Insurance Corporation. Accredited Health Facilities. updated 2026. https://www.philhealth.gov.ph/partners/providers/facilities/accredited/

[16] World Health Organization. Digital Adaptation Kit for Antenatal Care. 2021. https://www.who.int/publications/i/item/9789240020306

# Appendix A. Validation Evidence Register

| Evidence item | Status to record |
| --- | --- |
| Clinical rule review | Reviewer, date, version, comments, approval. |
| Automated test report | Commit/version, environment, pass/fail, defects. |
| Model data sheet | Source, population, outcome, features, missingness, limitations. |
| Model validation report | Split, metrics, CIs, calibration, subgroups, errors. |
| External/local validation | Site, period, sample, protocol, results. |
| Human-factors report | Users, tasks, scenarios, errors, comprehension, changes. |
| Referral readiness | Facilities, capabilities, contacts, transport, verification date. |
| Privacy/ethics approval | Institution, reference, scope, conditions. |
| Pilot protocol | Phase, endpoints, stopping rules, oversight. |
| Safety incident log | Event, severity, cause, action, closure. |
| Monitoring dashboard | Metrics, thresholds, reviewer, review frequency. |
| Release decision | Approvers, evidence reviewed, conditions, rollback plan. |

# Appendix B. Claim Restrictions

| Permitted claim after current work | Claim requiring additional evidence |
| --- | --- |
| The documented controller logic passed specified software tests. | The CDSS is clinically safe and effective. |
| The deterministic rules are based on cited guidelines and project safety decisions. | The rules diagnose the listed pregnancy conditions. |
| The model achieved specified internal test metrics, if transparently reported. | The model generalizes to Philippine lying-in patients. |
| The interface provides explainable reasons and recommendations. | Users understand and act correctly in real clinical care. |
| The referral workflow exists in the application. | Referrals are timely, completed, and improve outcomes. |
| The prototype supports supervised research and evaluation. | The system can replace professional judgment or operate autonomously. |
| The system may help standardize documentation and flag risk. | The system reduces maternal or neonatal mortality. |

# Appendix C. Local Site Readiness Checklist

Licensed clinic scope and staff roles documented.

Named supervising physician/obstetric clinical reviewer.

Current list of receiving hospitals and verified capabilities.

Emergency contact and transport pathway tested.

Stable device, power, backup, local network, and secure storage available.

Python inference service health check and manual fallback available.

Privacy notice, role permissions, password policy, and incident process approved.

Paper/downtime workflow and later reconciliation tested.

Midwife training and competency assessment completed.

Patient-facing explanation and language review completed.

Silent-pilot and supervised-pilot protocols approved.

Monitoring owner, frequency, stopping rules, and escalation authority assigned.